RISK ASSESSMENT

Dental Handpiece & Aerosol Management

Reference: BS-RA-DENT-007

Assessment Date: [Insert Date]

Assessor: [Insert Name]

Practice: [Insert Practice Name & Address]

Review Frequency: Annual (or following any incident/significant change)

1. SCOPE & PURPOSE

This risk assessment covers the risks associated with the use of dental handpieces (high-speed air turbines, slow-speed contra-angles) and ultrasonic scalers, specifically the generation of aerosols and splatter containing microorganisms, blood, and debris during clinical procedures.

2. RELEVANT LEGISLATION & GUIDANCE

- Health and Safety at Work etc. Act 1974 (HSWA) – Sections 2 & 3

- Control of Substances Hazardous to Health Regulations 2002 (COSHH) (SI 2002/2677)

- Management of Health and Safety at Work Regulations 1999 (SI 1999/3242)

- Personal Protective Equipment at Work Regulations 1992 (SI 1992/2966)

- Provision and Use of Work Equipment Regulations 1998 (PUWER) (SI 1998/2306)

- HTM 01-05: Decontamination in Primary Care Dental Practices (2013)

- HTM 03-01: Specialised Ventilation for Healthcare Premises (ventilation standards)

- SDCEP Mitigation of Aerosol Generating Procedures in Dentistry: A Rapid Review (2020, updated)

- CQC Fundamental Standards – Regulation 12 & Regulation 15

3. RISK RATING METHODOLOGY

Risk Rating = Likelihood (L) × Severity (S). Likelihood: 1 (Rare) to 5 (Almost Certain). Severity: 1 (Negligible) to 5 (Catastrophic).

| 1–3 (LOW) | Acceptable risk, monitor | 8–12 (MEDIUM) | Reduce risk where practicable |
| --- | --- | --- | --- |
| 4–6 (LOW-MED) | Monitor, review controls | 15–25 (HIGH) | Immediate action required |

4. RISK ASSESSMENT TABLE

| Hazard | Risk | Who at Risk | L | S | Rating | Control Measures | L | S | Residual |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Aerosol-borne pathogen transmission | Respiratory infection (including COVID-19, TB, influenza) | Staff/Patient | 3 | 4 | 12 (MEDIUM) | Pre-procedural antiseptic mouthwash (0.2% chlorhexidine or 1% hydrogen peroxide); high-volume aspiration (HVE) with wide-bore tip; rubber dam isolation where possible; FFP3 mask for aerosol-generating procedures (AGPs); surgery ventilation per HTM 03-01 | 1 | 3 | 3 (LOW) |
| Ocular splash/contamination | Eye infection, chemical splash injury | Staff/Patient | 3 | 3 | 9 (MEDIUM) | Full-face visor or safety goggles with side shields for operator and assistant; patient safety glasses provided; decontamination of eyewear between patients | 1 | 2 | 2 (LOW) |
| Handpiece aspiration contamination (suck-back) | Cross-contamination between patients via handpiece waterlines | Patient | 2 | 4 | 8 (MEDIUM) | Anti-retraction valves on all handpieces; handpieces run for 20–30 seconds between patients to flush waterlines; handpieces sterilised per HTM 01-05 after each patient use; DUWL treatment per RA-DENT-001 | 1 | 3 | 3 (LOW) |
| Hearing damage from high-speed handpiece noise | Noise-induced hearing loss (occupational) | Staff | 3 | 3 | 9 (MEDIUM) | Handpieces maintained per PUWER (worn bearings increase noise); noise levels below 80dB(A) daily exposure per Control of Noise at Work Regulations 2005; hearing protection available; health surveillance if exposure exceeds lower action value | 1 | 2 | 2 (LOW) |
| Fallow time non-compliance | Elevated aerosol exposure for next patient/staff | Staff/Patient | 3 | 3 | 9 (MEDIUM) | Fallow period per current SDCEP guidance (dependent on ventilation ACH rate); surgery ventilation verified (minimum 10 ACH recommended for AGPs); air purification (HEPA filtration) supplementary where ventilation insufficient; fallow timer used | 1 | 2 | 2 (LOW) |
| Handpiece mechanical failure during use | Bur detachment, laceration, aspiration of fragment | Patient | 2 | 4 | 8 (MEDIUM) | Pre-use check of bur seating; handpieces serviced per manufacturer schedule and PUWER; defective handpieces removed from service; bur inventory system (checking integrity before use); rubber dam as secondary protection | 1 | 3 | 3 (LOW) |

5. ADDITIONAL CONTROL MEASURES & EMERGENCY PROCEDURES

- Surgery ventilation assessed and verified (ACH rate measured) per SDCEP/HTM 03-01; results documented and reviewed annually.

- FFP3 masks fit-tested for all clinical staff performing AGPs; fit test records retained; repeat testing annually or when mask brand changes.

- Pre-procedural mouthwash protocol displayed in each surgery; patient compliance documented in clinical notes.

- High-volume evacuator (HVE) suction tips single-use or decontaminated per HTM 01-05; suction system serviced quarterly.

- AGP log maintained: date, surgery, procedure type, ventilation status, fallow time observed – available for CQC inspection.

6. MONITORING & REVIEW

This risk assessment shall be reviewed annually, or sooner if there is a significant change in working practices, an incident occurs, or new legislation/guidance is published. All staff must be made aware of the contents of this document and sign to confirm understanding.

7. SIGN-OFF

I confirm that this risk assessment has been carried out in accordance with relevant legislation and that the control measures identified are suitable and sufficient.

| Assessed By: [Insert Name] | Position: [Insert Name] | Date: [Insert Name] | Review Date: [Insert Name] |
| --- | --- | --- | --- |

© Barker Scott. All rights reserved. This document is confidential and intended for the sole use of the named dental practice.