RISK ASSESSMENT

Radiation Protection & Dental Radiography

Reference: BS-RA-DENT-002

Assessment Date: [Insert Date]

Assessor: [Insert Name]

Practice: [Insert Practice Name & Address]

Review Frequency: Annual (or following any incident/significant change)

1. SCOPE & PURPOSE

This risk assessment covers all dental radiographic procedures including intraoral periapical, bitewing, and occlusal radiographs, as well as extraoral panoramic (OPG) and cone beam computed tomography (CBCT) imaging, addressing radiation exposure risks to patients, staff, and members of the public.

2. RELEVANT LEGISLATION & GUIDANCE

- Ionising Radiations Regulations 2017 (IRR17) (SI 2017/1075)

- Ionising Radiation (Medical Exposure) Regulations 2017 (IR(ME)R 2017) (SI 2017/1322)

- Health and Safety at Work etc. Act 1974 (HSWA) – Sections 2 & 3

- Management of Health and Safety at Work Regulations 1999 (SI 1999/3242)

- Guidance Notes for Dental Practitioners on the Safe Use of X-Ray Equipment (PHE, 2020)

- British Dental Association (BDA) Advice Sheet: Radiation Protection in Dental Practice

- Selection Criteria for Dental Radiography (Faculty of General Dental Practice UK, 3rd Ed.)

- CQC Fundamental Standards – Regulation 12 (Safe Care and Treatment)

- Environmental Permitting (England and Wales) Regulations 2016 – disposal of X-ray equipment

3. RISK RATING METHODOLOGY

Risk Rating = Likelihood (L) × Severity (S). Likelihood: 1 (Rare) to 5 (Almost Certain). Severity: 1 (Negligible) to 5 (Catastrophic).

| 1–3 (LOW) | Acceptable risk, monitor | 8–12 (MEDIUM) | Reduce risk where practicable |
| --- | --- | --- | --- |
| 4–6 (LOW-MED) | Monitor, review controls | 15–25 (HIGH) | Immediate action required |

4. RISK ASSESSMENT TABLE

| Hazard | Risk | Who at Risk | L | S | Rating | Control Measures | L | S | Residual |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Unnecessary patient radiation exposure | Cumulative radiation dose, increased cancer risk | Patient | 3 | 4 | 12 (MEDIUM) | IR(ME)R 2017 compliance: justification by IRMER practitioner before each exposure; FGDP selection criteria applied; previous radiographs reviewed; dose recorded in patient notes; rectangular collimation and film holders to reduce scatter | 1 | 4 | 4 (LOW-MED) |
| Staff radiation exposure (primary/scatter) | Cumulative radiation dose, increased cancer risk | Staff | 2 | 4 | 8 (MEDIUM) | IRR17 Local Rules displayed in each X-ray area; controlled area designated with warning signs; operator stands behind protective barrier or >2m from tube at 90–135°; no staff member holds film/sensor; RPA and RPS appointed | 1 | 3 | 3 (LOW) |
| Public exposure from scatter radiation | Uncontrolled radiation dose to visitors/adjacent rooms | Public | 2 | 3 | 6 (LOW-MED) | Critical examination of new equipment by RPA before first clinical use; controlled area extends to 1.5m from X-ray tube head; warning lights operational; adjacent room shielding assessed by RPA; door closed during exposure | 1 | 2 | 2 (LOW) |
| Equipment malfunction (excess radiation output) | Over-exposure to patient and staff | Patient/Staff | 2 | 4 | 8 (MEDIUM) | 3-yearly equipment testing per IRR17 by Radiation Protection Adviser; annual quality assurance programme (image quality, exposure reproducibility); documented fault-reporting procedure; equipment maintenance log | 1 | 3 | 3 (LOW) |
| Failure to identify radiographic findings | Missed pathology, delayed diagnosis | Patient | 2 | 4 | 8 (MEDIUM) | All radiographs reported and findings documented in clinical notes; radiograph quality graded (1–3 per NRPB system); CPD in radiographic interpretation; peer review/audit of reporting | 1 | 3 | 3 (LOW) |
| Inadequate pregnancy screening | Foetal radiation exposure | Patient (pregnant) | 2 | 4 | 8 (MEDIUM) | Pregnancy status checked verbally and documented before X-ray; lead apron for abdominal shielding if clinically indicated; justification particularly rigorous for pregnant patients; signage in waiting area | 1 | 3 | 3 (LOW) |

5. ADDITIONAL CONTROL MEASURES & EMERGENCY PROCEDURES

- Radiation Protection Adviser (RPA) appointed and consulted at least 3-yearly (or when new equipment installed); Radiation Protection Supervisor (RPS) designated from dental team.

- Local Rules reviewed annually and signed by all staff; copies held in each X-ray surgery.

- IRMER roles documented: Referrer, Practitioner (Justifier), Operator – all staff acting in these roles hold appropriate training certificates.

- Quality assurance programme: monthly test exposures on phantom, annual dose area product (DAP) audit, image reject analysis.

- Significant accidental or unintended exposures reported to CQC per IR(ME)R 2017 Regulation 8; internal investigation and root cause analysis completed.

6. MONITORING & REVIEW

This risk assessment shall be reviewed annually, or sooner if there is a significant change in working practices, an incident occurs, or new legislation/guidance is published. All staff must be made aware of the contents of this document and sign to confirm understanding.

7. SIGN-OFF

I confirm that this risk assessment has been carried out in accordance with relevant legislation and that the control measures identified are suitable and sufficient.

| Assessed By: [Insert Name] | Position: [Insert Name] | Date: [Insert Name] | Review Date: [Insert Name] |
| --- | --- | --- | --- |

© Barker Scott. All rights reserved. This document is confidential and intended for the sole use of the named dental practice.